Molokai

Dx Internal Reproducibility Test Plan

Company Confidential

Revision History

| Version | Change# | Description of Change | Author | Date |
| --- | --- | --- | --- | --- |
| 1.0 | DCRUS03299 | Initial Release | T.Chandra, SH. Lim | 12-Apr-12 |
| 01.00 | DCO1051659 | Transferred from SAP 4477680 to Agile PPM (PCP0005347). Changed document P/N to PCP0005347 inside header. No content change inside main document. | Celine Roux | 24-Jul-13 |
| 02.00 | DCO1052795 | Update SAP PN to Agile PN, no change in document content | Celine Roux | 14-Aug-13 |

Table of Contents

1	Purpose	5

2	Scope	5

2.1	instruments, software and consumables	5

2.2	Veriti thermal cycler	6

3	Definitions	6

4	References	6

5	Materials	7

6	Requirements/Specifications to Be Tested	7

7	Run Matrix	7

7.1	Instruments and Test Methods	7

7.2	Instrument Maintenance	8

7.3	Spatial and Spectral calibrations and performance checks	8

7.4	Maintenance Wizards	8

7.5	Data Store and Run Records	8

7.6	Data Analysis	9

7.7	Defect Tracking	9

8	Specific Description of Tests to Be Performed	9

8.1	Study 1 – Testing Three (3) SeCore kits with one consumable lot and Three (3) UCLA HLA samples	9

8.1.1	Study 1 Instrument Setup	11

8.1.2	Study 1 DNA samples	11

8.1.3	Study 1 Plate Setup	11

8.1.4	Study 1 Run Setup	13

8.1.5	Study 1 Data Analysis	13

8.1.5.1	Genotype concordance	13

8.1.5.2	Signal	14

8.1.5.3	Noise to signal	14

8.2	Study 2 – Testing Three (3) Consumable Lots with One (1) SeCore kit and Three (3) UCLA HLA samples	14

8.2.1	Study 2 Run Setup	15

8.2.2	Study 2 Data Analysis	16

8.2.2.1	Genotype concordance	16

8.2.2.2	Signal	16

8.2.2.3	Noise to signal	16

Schedule	16

9	Approval	17

10	Appendix	18

10.1	Class I (Locus B) HLA samples	18

10.2	Class II (Locus DRB1) HLA samples	18

# Purpose

The purpose of Dx Internal Reproducibility testing is to confirm that the Applied Biosystems 3500(xl) Genetic Analyzer Dx system meets concordance reproducibility requirements for multiple  Dx consumable (CE) lots and SeCore kit lots. The  Dx 3500 system includes the 3500 Dx instrument,  Dx consumables,  Dx Data Collection Software and uTYPE Software and SeCore kits.

# Scope

## instruments, software and consumables

Dx Extended System Verification testing will be performed on one 24-capillary Applied Biosystems  3500 Genetic Analyzer Dx Evaluation unit.  The B (representative Class I locus) and DRB1 (representative Class II locus) SeCore kits will each be run on the instrument.

Two studies will be performed.  Both studies will be performed using Data Collection Software and uTYPE Software.

The first study will be performed using one (1) lot of Dx CE instrument consumables, three (3) lots of SeCore kits and one (1) operator over five (5) non-consecutive days.

The second study will be performed using three (3) lots of Dx CE instrument consumables, one (1) lot of SeCore kits and one (1) operator over five (5) non-consecutive days.

Three (3) lots of a representative SeCore GSSP kit for Locus B and Locus DRB1 will also be evaluated in the first study. Each SeCore kit lot will be matched with a unique GSSP kit lot. For Locus B the Z4 GSSP will be utilized and for Locus DRB1 the Codon 86TG GSSP will be utilized.

Similarly, a single lot of a representative SeCore GSSP kit (Z4 for Locus B and Codon 86TG for Locus DRB1) for Locus B and Locus DRB1 will be evaluated in the second study.

Table 2.1.1 lists the CE consumables and Table 2.1.2 lists the SeCore kits that will be used during testing for both studies. Testing will be performed on the  Data Collection Dx workflow.

Table 2.1.1 – CE instrument consumables

Hi-Di Formamide

ABC (Anode Buffer Container)

CBC (Cathode Buffer Container)

POP6 polymer

Conditioning pouch

Table 2.1.2 – SeCore and GSSP kits

SECORE LOC B SEQ KIT

SECORE LOC DRB1 SEQ KIT

Z4 (LOC B GSSP)

Codon 86TG (LOC DRB1 GSSP)

## Veriti thermal cycler

The Veriti RUO system will be used for testing.

# Definitions

Table 3.1 – Definitions

| Term | Definition |
| --- | --- |
| Cap | Capillary in array |
| DCR | Document Change Request |
| HiDi | Highly Deionized (brand name) |
| PRD | Product Requirements Document |
| PSD | Product Specifications Document |

# References

Table 4.1 – References

| Part Number | Reference |
| --- | --- |
| PCP0001043 | Consolidated Test Plan |
| PCP0001016 | Hardware PSD Appendix Table A-3 |
| VAL0001011 | SeCore PRDPSD trace matrix |
| MAN0003267 | SeCore and GSSP kits protocol |
| 4457397 | 3500 Dx User Guide |
| MAN0003385TBD | uTYPE User Guide |

# Materials

Materials listed in Table 5.1 with an (*) are Dx labeled materials.

Table 5.1 - Materials

*Hi-Di Formamide

*POP-6 Polymer 960 Samples

*Cathode Buffer Container

*Anode Buffer Container

*Conditioning Reagent

*Capillary Array (24-cap_50cm)

*96-well retainer & base standard

*Septa -96-well

*Septa Cathode Buffer Container

96-well plates

EtOH

# Requirements/Specifications to Be Tested

Table 6.1 – Requirements/Specifications to be tested

| Requirement/Specification | Metric | Passing Criterion |
| --- | --- | --- |
| SeCore | ≥ 95% Genotyping concordance at 95% confidence level | Lower Tolerance Limit of: Lower 95% confidence limit for Genotype concordance ≥ 95% for each locus evaluated |
| SeCore PSD-13.1 | The average signal of standard sequences will be ≥300. The average signal of group specific sequencing primers (GSSPs) sequences will be ≥100. | Lower Tolerance Limit of: Signal per base ≥ 300 (100 for GSSP) for ≥95% of all samples tested |
| SeCore PSD-12.1 | The average background will not exceed 8% | Upper Tolerance Limit of: %Noise:Signal < 8% for ≥ 95% of all Samples tested |

# Run Matrix

## Instruments and Test Methods

The instrument that will be used during testing is listed in Table 7.1.1. Testing will be performed in the Dx workflow of the software.

Table 7.1.1 – Instruments for testing

| Instrument Serial# | # of Capillaries |
| --- | --- |
| 23D303-20 | 24 |

## Instrument Maintenance

The instrument will be maintained per software notifications. Instruments will be pre-heated before each day’s run for 30 minutes.

## Spatial and Spectral calibrations and performance checks

.csv files for spatial and spectral calibrations, Performance check reports will be exported and saved.

## Maintenance Wizards

Wizard usage will be recorded in wizard run notes.

## Data Store and Run Records

Run notes will be filled out by the operator and reviewed by QA. Primary analysis will be done by the operator when data is reviewed and edited in uTYPE software. Statistical analysis will be done on the data afterwards and result will be included in the final test report. Section 7.6 Data Analysis lists all the applications tools and software to perform primary analysis (uTYPE) and statistical analysis.

.ab1 files and uTYPE reports generated by the operator will be stored in the project specific database.

## Data Analysis

The following applications will be needed for data analysis:

Table 7.5.1 – Data analysis applications

| Tools | Version |
| --- | --- |
| SeqA | 5.4 |
| Minitab | 7 |
| Tolerance Interval Script/Matlab | 1.1 |
| RIO Tool | 1.0 |
| uTYPE | 1.0 |
| uTYPE Library Release | July 2010 |

## Defect Tracking

Samples or files that fail amplification or sequencing will be repeated.  Samples that fail repeat testing due to the same issue will be documented and an investigation into the cause of failure will be initiated.  Samples that fail repeat testing due to a different cause may be repeated again.  Refer to the Troubleshooting section for general technical issues experienced during the study.

Instrument and consumable defects encountered during testing will be recorded in the instrument and consumables TeamTrack database.

# Specific Description of Tests to Be Performed

## Study 1 – Testing Three (3) SeCore kits with one consumable lot and Three (3) UCLA HLA samples

The SeCore B kit and the SeCore DRB1 kit will each be tested on one 24-cap  instrument. Pre-heat the instrument before the run is performed. Perform instrument maintenance per schedule. uTYPE software will be used to analyze results.  96-well plates will be used for all the runs. Each SeCore kit—Class I (B-locus) and Class II (DRB1)--will be run with 3 samples from the UCLA standard set of samples. Class I and Class II sample selection process started with GSSP mix selection since the GSSP mix is the limiting determinant.   GSSP selection was based on frequency in conjunction with sample availability in the DNA sample vendor catalog.  The GSSP positive samples were further refined to cover the maximum number of allele groups from 3 samples where possible.

Three (3) SeCore kit lots and one (1) instrument consumable lot will be integrated for testing.  Three samples per kit will be run.  Runs for each kit will be performed with one (1) operator over a time period of five (5) non-consecutive days. See Table 8.1.1 for run matrix. 75 total genotyping events per locus per SeCore kit lot will be obtained (3 samples/day x 5 reps/sample x 5 test days).

Sample Size Determination

For each locus, the concordance rate will be calculated and the corresponding exact 95% confidence interval will be calculated by the Clopper-Pearson method.  The lower confidence limit will be compared to the target value of approximately 0.95.  If the lower confidence limit is higher than 0.95, then it is concluded that the concordance rate is at least 0.95.  The minimum sample size (i.e.) genotyping results) needed to assess 95% concordance with 95% confidence is ln(0.05)/ln(0.95) = 59.

For each locus of interest, the concordance rate will be calculated and the corresponding exact 95% confidence interval will be calculated by the Clopper-Pearson method.  The lower confidence limit will be compared to the target value of 0.95.  If the lower confidence limit is higher than 0.95, then it is concluded that the concordance rate is at least 0.95.

Table 8.1.1 – Study 1 run matrix

| Operator | Inst. SN# | # of Caps | CE consumable Lot# | SeCore/GSSP kit lot# | SeCore kit | # of Class I Samples | SeCore kit | # of Class II Samples |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 23D303-20 | 24 | 1 | 1 | B | 3 | DRB1 | 3 |
| 1 | 23D303-20 | 24 | 1 | 2 | B | 3 | DRB1 | 3 |
| 1 | 23D303-20 | 24 | 1 | 3 | B | 3 | DRB1 | 3 |

### Study 1 Instrument Setup

- In the Dx workflow, install each instrument with a new capillary array.

- Install POP6 polymer, ABC and CBC according to instrument consumable lot#.

- Perform spatial calibration. Perform E spectral calibration and sequencing performance check with BDTv1.1 sequencing standard. Ensure that calibrations and performance checks are passing.

- Export spectral calibration .csv files and spatial, spectral and performance check reports.

### Study 1 DNA samples

UCLA DNA samples will be quantitated and diluted to 20 ng/ul using the Nandrop ND-1000.  The SeCore® product labeling indicates acceptable DNA concentration ranges for the samples as 15-30ng/μl but recommends using 20ng/μl. For this study, 20ng/ul will be tested.

### Study 1 Plate Setup

- Amplify DNA samples per Table 8.1.1 according to SeCore kit lot. Include molecular biology grade water as a negative control. Include sample C1-101 as a positive control for Class I kits and sample C2-101 as a positive control for Class II samples.

- Agarose gel check will be performed for each locus and sample.  A positive band in the negative control lane will require repeat of the assay set-up.  Agarose gel images will be saved.

- Create plates following the example plate layout in Table 8.1.2.1. Plate layouts will differ according to requirements of Study 1 (3 SeCore kit lots) and Study 2 (3 CE lots). Note: Record time at start of SeCore protocol and when plate is ready to be loaded onto instrument in run note.

Table 8.1.2.1 – Example plate layout for Study 1 and 2. MT = Empty well. Locus B (Class I) has 7 sequencing reactions per UCLA sample (2F, 2R, 3F, 3R, 4F, 4R and Z17) and Locus DRB1 (Class II) has 3 sequencing reactions (2F, 2R, and Codon 86TG) per UCLA sample. Test 5 replicates of each UCLA sample per day. 105 (=7 seq rxns x 5reps x 3samples) Class I sequencing reactions total for and 45 (=3 seq rxns x 5reps x 3amples) Class II sequencing reactions for each of 3 UCLA samples.

For Study 1: Repeat this plate layout for each of 3 SeCore reagent lots across 5 non-consecutive days using a single SeCore CE lot.

For Study 2: Repeat this plate layout for each of 3 SeCore CE lots across 5 non-consecutive days using one Secore kit lot.

| Plate 1 | 2F | 2R | 3F | 3R | 4F | 4R | Z17 | 2F | 2R | 86TG |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 |
| A | C1S1R1 | C1S1R1 | C1S1R1 | C1S1R1 | C1S1R1 | C1S1R1 | Z17 | C2S1R1 | C2S1R1 | 86TG | Pos Cont | Neg Cont |
| B | C1S2R1 | C1S2R1 | C1S2R1 | C1S2R1 | C1S2R1 | C1S2R1 | Z17 | C2S2R1 | C2S2R1 | 86TG | Pos Cont | Neg Cont |
| C | C1S3R1 | C1S3R1 | C1S3R1 | C1S3R1 | C1S3R1 | C1S3R1 | Z17 | C2S3R1 | C2S3R1 | 86TG | Pos Cont | Neg Cont |
| D | C1S1R2 | C1S1R2 | C1S1R2 | C1S1R2 | C1S1R2 | C1S1R2 | Z17 | C2S1R2 | C2S1R2 | 86TG | Pos Cont | Neg Cont |
| E | C1S2R2 | C1S2R2 | C1S2R2 | C1S2R2 | C1S2R2 | C1S2R2 | Z17 | C2S2R2 | C2S2R2 | 86TG | Pos Cont | Neg Cont |
| F | C1S3R2 | C1S3R2 | C1S3R2 | C1S3R2 | C1S3R2 | C1S3R2 | Z17 | C2S3R2 | C2S3R2 | 86TG | Pos Cont | Neg Cont |
| G | C1S1R3 | C1S1R3 | C1S1R3 | C1S1R3 | C1S1R3 | C1S1R3 | Z17 | C2S1R3 | C2S1R3 | 86TG | Pos Cont | Neg Cont |
| H | C1S2R3 | C1S2R3 | C1S2R3 | C1S2R3 | C1S2R3 | C1S2R3 | Z17 | C2S2R3 | C2S2R3 | 86TG | Pos Cont | Neg Cont |

C=UCLA Class (C1=Class I Locus B, C2=Class II Locus DRB1)

S=UCLA Sample number (S1 =Sample 1, S2 = Sample 2, S3 = Sample 3)

R=Replicate (R1=Replicates 1, R2=Replicate 2, R3=Replicate 3, R4=Replicate 4, R5 = Replicate 5)

| Plate 2 | 2F | 2R | 3F | 3R | 4F | 4R | Z17 | 2F | 2R | 86TG |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 |
| A | C1S3R3 | C1S3R3 | C1S3R3 | C1S3R3 | C1S3R3 | C1S3R3 | Z17 | C2S3R3 | C2S3R3 | 86TG | Pos Cont | Neg Cont |
| B | C1S1R4 | C1S1R4 | C1S1R4 | C1S1R4 | C1S1R4 | C1S1R4 | Z17 | C2S1R4 | C2S1R4 | 86TG | Pos Cont | Neg Cont |
| C | C1S2R4 | C1S2R4 | C1S2R4 | C1S2R4 | C1S2R4 | C1S2R4 | Z17 | C2S2R4 | C2S2R4 | 86TG | Pos Cont | Neg Cont |
| D | C1S3R4 | C1S3R4 | C1S3R4 | C1S3R4 | C1S3R4 | C1S3R4 | Z17 | C2S3R4 | C2S3R4 | 86TG | Pos Cont | Neg Cont |
| E | C1S1R5 | C1S1R5 | C1S1R5 | C1S1R5 | C1S1R5 | C1S1R5 | Z17 | C2S1R5 | C2S1R5 | 86TG | Pos Cont | Neg Cont |
| F | C1S2R5 | C1S2R5 | C1S2R5 | C1S2R5 | C1S2R5 | C1S2R5 | Z17 | C2S2R5 | C2S2R5 | 86TG | Pos Cont | Neg Cont |
| G | C1S3R5 | C1S3R5 | C1S3R5 | C1S3R5 | C1S3R5 | C1S3R5 | Z17 | C2S3R5 | C2S3R5 | 86TG | Pos Cont | Neg Cont |
| H | MT | MT | MT | MT | MT | MT | MT | MT | MT | MT | Pos Cont | Neg Cont |

C=UCLA Class (C1=Class I - Locus B, C2=Class II - Locus DRB1)

S=UCLA Sample number (S1 =Sample 1, S2 = Sample 2, S3 = Sample 3)

R=Replicate (R1=Replicates 1, R2=Replicate 2, R3=Replicate 3, R4=Replicate 4, R5 = Replicate 5)

### Study 1 Run Setup

- Pre-heat instrument. Run plate on each instrument per run matrix.

- Perform 1 injection of the RapidSeqPOP6 module for each SeCore kit sample.

- Load new consumables as required.

### Study 1 Data Analysis

#### Genotype concordance

- Import each sample into uTYPE software. Refer to uTYPE IFU.

- Perform edits as necessary.

- Save uTYPE report with allele assignment.

- Evaluate if the sample matches the known genotype. Record if the sample matches the known genotype.

- Calculate the % of samples that match the known genotype for each SeCore kit.

- Use the tolerance interval tool to calculate the lower tolerance of the %  concordance with 95% confidence for each SeCore kit type.

#### Signal

- Run the RIO tool for each sample

- Use the total signal and signal balance values to calculate the signal per base.

- For each SeCore kit, use the tolerance interval tool to calculate the lower tolerance of the signal per base with 95% confidence and 90% proportion.

- Run the Q Value report in uTYPE for each SeCore kit and each run.  Q Value report automatically calculates the average signal and noise to signal values for each sample, exon and direction.

#### Noise to signal

- Import each sample into uTYPE. Refer to uTYPE IFU.

- Use the Q-Report noise to signal values in the report for each sample and each run.

- For each SeCore kit, use the tolerance interval tool to calculate the upper tolerance of the signal per base with 95% confidence and 95% concordance.

## Study 2 – Testing Three (3) Consumable Lots with One (1) SeCore kit and Three (3) UCLA HLA samples

The SeCore B kit and the SeCore DRB1 kit will each be tested on one 24-cap instrument. Pre-heat the instrument before the run is performed. Perform instrument maintenance per schedule. uTYPE software will be used to analyze results.

Each SeCore kit—Class I (B-locus) and Class II (DRB1)--will be run with 3 samples from the UCLA standard set of samples. Class I and Class II sample selection process started with GSSP mix selection since the GSSP mix is the limiting determinant.   GSSP selection was based on frequency in conjunction with sample availability in the DNA sample vendor catalog.  The GSSP positive samples were further refined to cover the maximum number of allele groups from 3 samples where possible.

One (1) SeCore kit lot and three (3) instrument consumable lots will be integrated for testing. Three samples per kit will be run.  Runs for each kit will be performed with one (1) operator over a time period of five (5) non-consecutive days. 75 total genotyping events per locus per SeCore kit lot will be obtained. See Table 8.1.2.1 for plate layout and Table 8.1.2 for run matrix.

Sample Size Determination

For each locus, the concordance rate will be calculated and the corresponding exact 95% confidence interval will be calculated by the Clopper-Pearson method.  The lower confidence limit will be compared to the target value of approximately 0.95.  If the lower confidence limit is higher than 0.95, then it is concluded that the concordance rate is at least 0.95.  The minimum sample size (i.e.) genotyping results) needed to assess 95% concordance with 95% confidence is ln(0.05)/ln(0.95) = 59.

Table 8.1.2 – Study 2 run matrix.

| Operator | Inst.Name/SN# | # of Caps | CE consumable Lot# | SeCore/GSSP kit lot# | SeCore kit | # of Class I Samples | SeCore kit | # of Class II Samples |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 23D303-20 | 24 | 1 | 1 | B | 3 | DRB1 | 3 |
| 1 | 23D303-20 | 24 | 2 | 1 | B | 3 | DRB1 | 3 |
| 1 | 23D303-20 | 24 | 3 | 1 | B | 3 | DRB1 | 3 |

### Study 2 Run Setup

- Pre-heat instrument. Run plate on each instrument per run matrix.

- Perform 1 injection of the RapidSeqPOP6 module for each SeCore kit sample.

- Load new consumables as required.

### Study 2 Data Analysis

#### Genotype concordance

- Import each sample into uTYPE software. Refer to uTYPE IFU.

- Perform edits as necessary.

- Save uTYPE report with allele assignment.

- Evaluate if the sample matches the known genotype. Record if the sample matches the known genotype.

- Calculate the % of samples that match the known genotype for each unique SeCore kit/consumable lot combination.

- Use the tolerance interval tool to calculate the lower tolerance of the % matching samples with 95% confidence and 90% proportion for each SeCore kit type.

#### Signal

- Run the RIO tool for each sample

- Use the total signal and signal balance values to calculate the signal per base.

- For each unique SeCore kit/consumable lot combination, use the tolerance interval tool to calculate the lower tolerance of the signal per base with 95% confidence and 90% proportion.

- Run the Q Value report in uTYPE for each SeCore kit and each run.  Q Value report automatically calculates the average signal and noise to signal values for each sample, exon and direction.

#### Noise to signal

- Import each sample into uTYPE. Refer to uTYPE IFU.

- Use the Q-Report to calculate noise to signal values for each sample and run

- For each unique SeCore kit/consumable lot combination, use the tolerance interval tool to calculate the upper tolerance of the signal per base with 95% confidence and 90% proportion.

# Schedule

| Activity | Duration |
| --- | --- |
| Dx Extended Verification testing | 40 days |

# Approval

See Agile DCO for document approval.

# Appendix

## Class I (Locus B) HLA samples

The tables below lists the UCLA Class I HLA samples and controls to be run. The samples names listed below are masked names.

Table 10.1.1 – Class I samples

Class I

D134

D136

D137

## Class II (Locus DRB1) HLA samples

The tables below lists the UCLA Class II HLA samples and controls to be run. The samples names listed below are masked names.

Table 10.2.1 – Class II samples

Class II

F136

F109

F114